## New Attributes
Tag | Name | Keyword | VR | VM | Private Tag | Indent | Per Frame | Notes
(gggg,eee1) | PA Excitation Characteristics Sequence | PAExcitationCharacteristicsSequence | SQ | 1 | (3411, 1001) | 0 | 1
(gggg,eee2) | Excitation Spectral Width | ExcitationSpectralWidth | FL | 1 | (3451, 1002) | 1 | 1 | Randomized value for per-frame illustration
(gggg,eee3) | Excitation Energy | ExcitationEnergy | FL | 1 | (3451, 1003) | 1 | 1 | Randomized value for per-frame illustration
(gggg,eee4) | Excitation Pulse Duration | ExcitationPulseDuration | FL | 1 | (3451, 1004) | 1 | 1 | Randomized value for per-frame illustration
(gggg,ee94) | Excitation Wavelength Sequence | ExcitationWavelengthSequence | SQ | 1 | (3401, 1094) | 0 | 0
(gggg,eee5) | Excitation Wavelength | ExcitationWavelength | FL | 1 | (3441, 1005) | 1 | 0
(gggg,ee92) | Illumination Translation Flag | IlluminationTranslationFlag | CS | 1 | (3401, 1092) | 0 | 0
(gggg,eee6) | Illumination Type Code Sequence | IlluminationTypeCodeSequence | SQ | 1 | (3401, 1006) | 0 | 0 | Followed by Illumination Type Code Sequence table
(gggg,ee99) | Acoustic Coupling Medium Flag | AcousticCouplingMediumFlag | CS | 1 | (3401, 1099) | 0 | 0
(gggg,eee7) | Acoustic Coupling Medium Code Sequence | AcousticCouplingMediumCodeSequence | SQ | 1 | (3401, 1007) | 0 | 0 | Followed by Acoustic Coupling Medium Code Sequence table
(gggg,eee8) | Coupling Medium Temperature | CouplingMediumTemperature | FL | 1 | (3401, 1008) | 0 | 0
(gggg,ee17) | Transducer Response Sequence | TransducerResponseSequence | SQ | 1 | (3401, 1017) | 0 | 0
(gggg,ee98) | Center Frequency | CenterFrequency | UL | 1 | (3431, 1098) | 1 | 0
(gggg,ee97) | Fractional Bandwidth | FractionalBandwidth | UL | 1 | (3431, 1097) | 1 | 0
(gggg,ee96) | Lower Cutoff Frequency | LowerCutoffFrequency | UL | 1 | (3431, 1096) | 1 | 0
(gggg,ee95) | Upper Cutoff Frequency | UpperCutoffFrequency | UL | 1 | (3431, 1095) | 1 | 0
(gggg,ee10) | Transducer Technology Sequence | TransducerTechnologySequence | SQ | 1 | (3401, 1010) | 0 | 0 | Followed by Ultrasound Transducer Technology Code Sequence
(gggg,ee14) | Sound Speed Correction Mechanism Code Sequence | SoundSpeedCorrectionMechanismCodeSequence | SQ | 1 | (3401, 1014) | 0 | 0 | Followed by Speed of Sound Correction Mechanisms Code Sequence, Referenced Image Sequence (not required for first example)
(gggg,ee15) | Object Sound Speed | ObjectSoundSpeed | FL | 1 | (3421, 1015) | 1 | 0
(gggg,ee1a) | Coupling Medium Sound Speed | CouplingMediumSoundSpeed | FL | 1 | (3421, 101a) | 1 | 0 | Not required for first example
(gggg,ee93) | PA Dimension Index ID | PADimensionIndexID | FL | 1 | (3401, 1093) | 0 | 0
(gggg,a001) | PA Image Frame Type Sequence | PAImageFrameTypeSequence | SQ | 1 | (3461, 10a1) | 0 | 1 | Same as USImageDescriptionSequence - leave as US for now to help viewers

## New Codes
Coding Scheme Designator | Code Value | Code Meaning | Temporary Code
DCM | XXXXX1 | Single-side illumination | 103401
DCM | XXXXX2 | Dual-side illumination | 103402
DCM | XXXXX3 | Multi-side illumination | 103403
DCM | XXXXX4 | Through-transducer illumination | 103404
DCM | XXXX13 | Piezocomposite Transducer | 103413
DCM | XXXX14 | MEMS-based Transducer | 103414
DCM | XXXX15 | Interferometric Transducer | 103415
DCM | XXXX16 | Uniform Speed of Sound Correction | 103416
DCM | XXXX17 | Dual Speed of Sound Correction | 103417
DCM | XXXX18 | Speed of Sound Map Correction | 103418
DCM | XXXXXA | Planar matrix ultrasound transducer geometry | 103420
DCM | XXXXXB | Hemispherical ultrasound transducer geometry | 103421

## Existing Attributes
Name | Tag | Indent | Per Frame | Notes
Acquisition DateTime | (0008,002A) | 0 | 0
Position Measuring Device Used | (0018,980C) | 0 | 0
Lossy Image Compression | (0028,2110) | 0 | 0
Lossy Image Compression Ratio | (0028,2112) | 0 | 0
Lossy Image Compression Method | (0028,2114) | 0 | 0
Presentation LUT Shape | (2050,0020) | 0 | 0
Event Timer Sequence | (0008,2133) | 0 | 0 | Not required for first example
Event Time Offset | (0008,2134) | 1 | 0
Event Code Sequence | (0008,2135) | 1 | 0 | Followed by Protocol Interval Events Code Sequence
Event Timer Name(s) | (0008,2132) | 1 | 0
Transducer Geometry Code Sequence | (0018,980D) | 0 | 0 | Followed by Ultrasound Transducer Geometry Code Sequence
Reconstruction Algorithm Sequence | (0018,993D) | 0 | 0 | Followed by the Algorithm Identification Table
Frame Type | (0008,9007) | 1 | 1 | Will be in PAImageFrameTypeSequence -same as USImageDescriptionSequence - currently included as US for now to help viewers
Volumetric Properties | (0008,9206) | 1 | 1
Volume Based Calculation Technique | (0008,9207) | 1 | 1
SOP Class UID | (0008,0016) | 0 | 0 | PhotoacousticImageStorage, temporary value of 1.2.840.10008.5.1.2.3.45

## DCMTk-private.dic
(3441,"WG-34 PA Excitation WL",1005) | FL | ExcitationWavelength | 1 | PrivateTag
(3401,"WG-34 PA Proposed Tags",1006) | SQ | IlluminationTypeCodeSequence | 1 | PrivateTag
(3401,"WG-34 PA Proposed Tags",1007) | SQ | AcousticCouplingMediumCodeSequence | 1 | PrivateTag
(3401,"WG-34 PA Proposed Tags",1008) | FL | CouplingMediumTemperature | 1 | PrivateTag
(3401,"WG-34 PA Proposed Tags",1010) | SQ | TransducerTechnologySequence | 1 | PrivateTag
(3401,"WG-34 PA Proposed Tags",1014) | SQ | SoundSpeedCorrectionMechanismCodeSequence | 1 | PrivateTag
(3421,"WG-34 PA Sound Speed",1015) | FL | ObjectSoundSpeed | 1 | PrivateTag
(3401,"WG-34 PA Proposed Tags",1017) | SQ | TransducerResponseSequence | 1 | PrivateTag
(3421,"WG-34 PA Sound Speed",101A) | FL | CouplingMediumSoundSpeed | 1 | PrivateTag
(3401,"WG-34 PA Proposed Tags",1092) | CS | IlluminationTranslationFlag | 1 | PrivateTag
(3401,"WG-34 PA Proposed Tags",1093) | FL | PADimensionIndexID | 1 | PrivateTag
(3401,"WG-34 PA Proposed Tags",1094) | SQ | ExcitationWavelengthSequence | 1 | PrivateTag
(3431,"WG-34 PA Transducer Response",1095) | UL | UpperCutoffFrequency | 1 | PrivateTag
(3431,"WG-34 PA Transducer Response",1096) | UL | LowerCutoffFrequency | 1 | PrivateTag
(3431,"WG-34 PA Transducer Response",1097) | UL | FractionalBandwidth | 1 | PrivateTag
(3431,"WG-34 PA Transducer Response",1098) | UL | CenterFrequency | 1 | PrivateTag
(3401,"WG-34 PA Proposed Tags",1099) | CS | AcousticCouplingMediumFlag | 1 | PrivateTag
(3411,"WG-34 PA Per-Frame Seq",1001) | SQ | PAExcitationCharacteristicsSequence | 1 | PrivateTag
(3451,"WG-34 PA Per-Frame Tags",1002) | FL | ExcitationSpectralWidth | 1 | PrivateTag
(3451,"WG-34 PA Per-Frame Tags",1003) | FL | ExcitationEnergy | 1 | PrivateTag
(3451,"WG-34 PA Per-Frame Tags",1004) | FL | ExcitationPulseDuration | 1 | PrivateTag
(3461,"WG-34 PA Image Frame Seq",10a1) | SQ | PAImageFrameTypeSequence | 1 | PrivateTag