## New Attributes
Tag | Name | Keyword | LB VR | FT VR | VM | LB Private Tag | FT Private Tag | Indent | Per Frame | Notes
(gggg,eee1) | Photoacoustic Excitation Characteristics Sequence | Photoacoustic​Excitation​Characteristics​Sequence | SQ | SQ | 1 | (3411, 1001) | (0018,9821) | 0 | 1
(gggg,eee2) | Excitation Spectral Width | ExcitationSpectralWidth | FL | FD | 1 | (3451, 1002) | (0018,9822) | 1 | 1 | Randomized value for per-frame illustration
(gggg,eee3) | Excitation Energy | ExcitationEnergy | FL | FD | 1 | (3451, 1003) | (0018,9823) | 1 | 1 | Randomized value for per-frame illustration
(gggg,eee4) | Excitation Pulse Duration | ExcitationPulseDuration | FL | FD | 1 | (3451, 1004) | (0018,9824) | 1 | 1 | Randomized value for per-frame illustration
(gggg,ee94) | Excitation Wavelength Sequence | ExcitationWavelengthSequence | SQ | SQ | 1 | (3401, 1094) | (0018,9825) | 0 | 0
(gggg,eee5) | Excitation Wavelength | ExcitationWavelength | FL | FD | 1 | (3441, 1005) | (0018,9826) | 1 | 0 | Excitation wavelength is found in two places - Photoacoustic Acquisition Parameters Module Attributes and in the Per-Frame Photoacoustic Excitation Characteristics Macro
(gggg,ee92) | Illumination Translation Flag | IlluminationTranslationFlag | CS | CS | 1 | (3401, 1092) | (0018,9828) | 0 | 0
(gggg,eee6) | Illumination Type Code Sequence | IlluminationTypeCodeSequence | SQ | CS | 1 | (3401, 1006) | (0022,0016) | 0 | 0 | Previously listed as a new attribute, but found to have an existing name/tag. Followed by Illumination Type Code Sequence table
(gggg,ee99) | Acoustic Coupling Medium Flag | AcousticCouplingMediumFlag | CS | CS | 1 | (3401, 1099) | (0018,9829) | 0 | 0
(gggg,eee7) | Acoustic Coupling Medium Code Sequence | AcousticCouplingMediumCodeSequence | SQ | SQ | 1 | (3401, 1007) | (0018,982A) | 0 | 0 | Followed by Acoustic Coupling Medium Code Sequence table
(gggg,eee8) | Acoustic Coupling Medium Temperature | Acoustic​Coupling​Medium​Temperature | FL | FD | 1 | (3401, 1008) | (0018,982B) | 0 | 0
(gggg,ee17) | Transducer Response Sequence | TransducerResponseSequence | SQ | SQ | 1 | (3401, 1017) | (0018,982C) | 0 | 0
(gggg,ee98) | Center Frequency | CenterFrequency | FL | FD | 1 | (3431, 1098) | (0018,982D) | 1 | 0
(gggg,ee97) | Fractional Bandwidth | FractionalBandwidth | FL | FD | 1 | (3431, 1097) | (0018,982E) | 1 | 0
(gggg,ee96) | Lower Cutoff Frequency | LowerCutoffFrequency | FL | FD | 1 | (3431, 1096) | (0018,982F) | 1 | 0
(gggg,ee95) | Upper Cutoff Frequency | UpperCutoffFrequency | FL | FD | 1 | (3431, 1095) | (0018,9830) | 1 | 0
(gggg,ee10) | Transducer Technology Sequence | TransducerTechnologySequence | SQ | SQ | 1 | (3401, 1010) | (0018,9831) | 0 | 0 | Followed by Ultrasound Transducer Technology Code Sequence
(gggg,ee14) | Sound Speed Correction Mechanism Code Sequence | SoundSpeedCorrectionMechanismCodeSequence | SQ | SQ | 1 | (3401, 1014) | (0018,9832) | 0 | 0 | Followed by Speed of Sound Correction Mechanisms Code Sequence, Referenced Image Sequence (not required for first example)
(gggg,ee15) | Object Sound Speed | ObjectSoundSpeed | FL | FD | 1 | (3421, 1015) | (0018,9833) | 1 | 0
(gggg,ee1a) | Acoustic Coupling Medium Sound Speed | Acoustic​Coupling​Medium​Sound​Speed | FL | FD | 1 | (3421, 101a) | (0018,9834) | 1 | 0 | Not required for first example
(gggg,a001) | Photoacoustic Image Frame Type Sequence | Photoacoustic​Image​Frame​Type​Sequence |  | SQ | 1 | (3461, 10a1) | (0018,9835) | 0 | 0 | Same structure as USImageDescriptionSequence - can leave as US for now to help viewers
| Image Data Type Code Sequence | Image​Data​Type​Code​Sequence | SQ | SQ | 1 |  | (0018,9836) | 1 | 1 | Added in final design. Below Image Data Type Sequence. Analogous to (0018,9808) Data Type for US, but PA is using codes instead of defined terms.
(gggg,ee93) | PA Reconstruction Index | PAReconstructionIndex | UL |  |  | (3401, 1093) |  |  |  | Dropped from the final design

## Data Dictionary Changes LB-FT
FT Tag | FT Name | FT Keyword | FT VR | FT VM | FT Indent | FT Per Frame | FT Shared | LB Name
(0018,9821) | Photoacoustic Excitation Characteristics Sequence | Photoacoustic​Excitation​Characteristics​Sequence | SQ | 1 | 0 | 1 | 1 | PA Excitation Characteristics Sequence
(0018,9822) | Excitation Spectral Width | Excitation​Spectral​Width | FD | 1 | 1 | 1 | 1 | Excitation Spectral Width
(0018,9823) | Excitation Energy | Excitation​Energy | FD | 1 | 1 | 1 | 1 | Excitation Energy
(0018,9824) | Excitation Pulse Duration | Excitation​Pulse​Duration | FD | 1 | 1 | 1 | 1 | Excitation Pulse Duration
(0018,9825) | Excitation Wavelength Sequence | Excitation​Wavelength​Sequence | SQ | 1 | 0 | 0 | 0 | Excitation Wavelength Sequence
(0018,9826) | Excitation Wavelength | Excitation​Wavelength | FD | 1 | 1 | 0 | 0 | Excitation Wavelength
(0018,9828) | Illumination Translation Flag | Illumination​Translation​Flag | CS | 1 | 0 | 0 | 0 | Illumination Translation Flag
(0022,0016) | Illumination Type Code Sequence | IlluminationTypeCodeSequence | CS | 1 | 0 | 0 | 0 | Illumination Type Code Sequence
(0018,9829) | Acoustic Coupling Medium Flag | Acoustic​Coupling​Medium​Flag | CS | 1 | 0 | 0 | 0 | Acoustic Coupling Medium Flag
(0018,982A) | Acoustic Coupling Medium Code Sequence | Acoustic​Coupling​Medium​Code​Sequence | SQ | 1 | 0 | 0 | 0 | Acoustic Coupling Medium Code Sequence
(0018,982B) | Acoustic Coupling Medium Temperature | Acoustic​Coupling​Medium​Temperature | FD | 1 | 0 | 0 | 0 | Coupling Medium Temperature
(0018,982C) | Transducer Response Sequence | Transducer​Response​Sequence | SQ | 1 | 0 | 0 | 0 | Transducer Response Sequence
(0018,982D) | Center Frequency | Center​Frequency | FD | 1 | 1 | 0 | 0 | Center Frequency
(0018,982E) | Fractional Bandwidth | Fractional​Bandwidth | FD | 1 | 1 | 0 | 0 | Fractional Bandwidth
(0018,982F) | Lower Cutoff Frequency | Lower​Cutoff​Frequency | FD | 1 | 1 | 0 | 0 | Lower Cutoff Frequency
(0018,9830) | Upper Cutoff Frequency | Upper​Cutoff​Frequency | FD | 1 | 1 | 0 | 0 | Upper Cutoff Frequency
(0018,9831) | Transducer Technology Sequence | Transducer​Technology​Sequence | SQ | 1 | 0 | 0 | 0 | Transducer Technology Sequence
(0018,9832) | Sound Speed Correction Mechanism Code Sequence | Sound​Speed​Correction​Mechanism​Code​Sequence | SQ | 1 | 0 | 0 | 0 | Sound Speed Correction Mechanism Code Sequence
(0018,9833) | Object Sound Speed | Object​Sound​Speed | FD | 1 | 1 | 0 | 0 | Object Sound Speed
(0018,9834) | Acoustic Coupling Medium Sound Speed | Acoustic​Coupling​Medium​Sound​Speed | FD | 1 | 1 | 0 | 0 | Coupling Medium Sound Speed
(0018,9835) | Photoacoustic Image Frame Type Sequence | Photoacoustic​Image​Frame​Type​Sequence | SQ | 1 | 0 | 0 | 1 | PA Reconstruction Index
(0018,9836) | Image Data Type Code Sequence | Image​Data​Type​Code​Sequence | SQ | 1 | 1 | 0 | 1 | PA Image Frame Type Sequence

## New Codes
Coding Scheme Designator | Code Value | Code Meaning | Temporary Code | FT Code
DCM | XXXXX1 | Single-side illumination | 103401 | 130810
DCM | XXXXX2 | Dual-side illumination | 103402 | 130811
DCM | XXXXX3 | Multi-side illumination | 103403 | 130812
DCM | XXXXX4 | Through-transducer illumination | 103404 | 130813
DCM | XXXXX5 | Interstitial illumination | 103405 | 130814
DCM | XXXX13 | Piezocomposite Transducer | 103413 | 130815
DCM | XXXX14 | MEMS-based Transducer | 103414 | 130816
DCM | XXXX15 | Interferometric Transducer | 103415 | 130817
DCM | XXXX16 | Uniform Speed of Sound Correction | 103416 | 130818
DCM | XXXX17 | Dual Speed of Sound Correction | 103417 | 130819
DCM | XXXX18 | Speed of Sound Map Correction | 103418 | 130820
DCM | XXXXXA | Planar matrix ultrasound transducer geometry | 103420 | 130808
DCM | XXXXXB | Hemispherical ultrasound transducer geometry | 103421 | 130809
DCM |  | Spherical Back Projection |  | 130821
DCM |  | Numerical Acoustic Model |  | 130822
DCM |  | Blood Oxygenation Level |  | 110819
DCM |  | Blood Deoxygenation Level |  | 130823
SCT |  | Hemoglobin |  | 38082009
DCM |  | Elasticity |  | 110830
DCM |  | Perfusion |  | 110831
SCT |  | Melanin |  | 59094002
SCT |  | Water |  | 11713004
SCT |  | Lipid |  | 70106000

## Existing Attributes
Name | Tag | Indent | Per Frame | Shared | Notes
Acquisition DateTime | (0008,002A) | 0 | 0
Position Measuring Device Used | (0018,980C) | 0 | 0
Lossy Image Compression | (0028,2110) | 0 | 0
Lossy Image Compression Ratio | (0028,2112) | 0 | 0
Lossy Image Compression Method | (0028,2114) | 0 | 0
Presentation LUT Shape | (2050,0020) | 0 | 0
Transducer Geometry Code Sequence | (0018,980D) | 0 | 0 |  | Followed by Ultrasound Transducer Geometry Code Sequence
Reconstruction Algorithm Sequence | (0018,993D) | 0 | 0 |  | Followed by the Algorithm Identification Table
Frame Type | (0008,9007) | 1 | 1 |  | Will be in PAImageFrameTypeSequence -same as USImageDescriptionSequence - currently included as US for now to help viewers
Volumetric Properties | (0008,9206) | 1 | 1
Volume Based Calculation Technique | (0008,9207) | 1 | 1
SOP Class UID | (0008,0016) | 0 | 0 |  | PhotoacousticImageStorage, temporary value of 1.2.840.10008.5.1.2.3.45 | Final Text SOP Class UID: 1.2.840.10008.5.1.4.1.1.6.3
Image Data Type Sequence | (0018,9807) | 0 | 0 | 1
Illumination Type Code Sequence | (0022,0016) | 0 | 0 | 0 | A previously existing name/tag. Followed by Illumination Type Code Sequence table